DRAFT – 12/30/00

Processing of surface meteorological data collected on the R/V Revelle winter SeaSoar cruise (Revelle3)

Introduction

This report describes the data return, sensor comparison, and basic processing of the surface meteorological data collected on the R/V Revelle winter SeaSoar cruise (Revelle3).  The final “best” basic 1-min time series represents a combination of data collected with the ship’s IMET system, WHOI ASIMET short- and long-wave radiation recorders, and SDSU units that measured wind speed and direction, air temperature, relative humidity, and air pressure.  The SDSU units recorded data every second while the IMET and ASIMET units used 1-min record intervals.

The 1-sec SDSU wind data were vectored averaged and then low-pass filtered and subsampled to obtain a 1-min time series.  When possible, comparisons of data collected with different sensor systems have been used to help determine which data should be used for scientific analysis.

The “best” 1-min time series are described in revel3_1m.txt and stored in the MATLAB 5 file revel3_1m.mat.  This 1-min data was then low-pass filtered with PL66TN (half amplitude period = 12 min) and subsampled to produce final 5-min time series for scientific analysis.  The 5-min data were used to estimate surface fluxes using the MATLAB Air-Sea Toolbox (Pawlowicz et al, 2000; Beardsley et al, 1998). The 5-min time series are described in revel3_5m.txt and stored in revel3_5m.mat.   These files are available at the anonymous ftp site ftp.whoi.edu/pub/users/rcbshare.

Wind Speed and Direction

The IMET wind sensor mounted on the bow mast returned 1-min wind speed for most of the cruise with a 4-day gap in the center of the cruise.  The direction sensor failed (returning only constant directions), so that the IMET vector average winds were bad.

C. Dorman mounted SDSU wind speed and direction sensor sets at two levels on the bow mast and on the port and starboard wings of the main mast.  These sensors returned 1-sec data that were then vector-averaged into 1-min apparent wind speed and direction for subsequent analysis.  Comparison of these winds suggested that the upwind sensor set on the main mast gave the best measurement (with no obvious sheltering effects, unlike the bow mast sensors).  Comparison of the upwind main mast wind speed with the IMET wind speed when the ship was steaming into the wind indicated good agreement, with no adjustment needed for the difference in sensor height.  Thus, the upwind main mast SDSU 1-min wind speed and direction data were used with a sensor height of 15 m in all analysis.

True winds were computed using the SDSU measured relative wind speed and direction, the ship’s heading as given by the gyro, and the motion of the ship as given by the P-code GPS.

Air Temperature

The IMET air temperature sensor on the bow mast returned good data for the entire cruise.  A SDSU air temperature sensor, mounted on the bow mast about 0.5 m below the IMET sensor, returned data for periods when the air temperature was above about –4(C.  This data, recorded at 0.25 sec, was low-pass filtered and subsampled to obtain 3687 temperature values coincident with the 5-min IMET series.  The two series were highly correlated (cc2 = 0.990), with a mean IMET minus SDSU difference and standard deviation of 0.28 ( 0.25(C, the IMET reading about 0.3 OC higher on average.  Linear regression gave the fit

ATSDSU  = a + b ATIMET ,

with a = -0.36 ( 0.02(C, b = 1.040 ( 0.003, and a standard deviation of ( 0.23(C about the fit (Figure 1).   The IMET air temperature data were used in subsequent analysis.

Figure 1. Regression of the IMET and SDSU 5-min air temperature data for

periods when the air temperature was above –4(C.

Relative Humidity

The IMET relative humidity sensor on the bow mast returned good data for the entire cruise.  A SDSU relative humidity sensor, mounted on the bow mast about 0.5 m below the IMET sensor, returned data that generally tracked the IMET data when the air temperature was above –4(C.  Following the same procedure used to compare air temperature, a set of SDSU 3687 relative humidity values coincident with the 5-min IMET series was obtained.  With the exception of several short periods when the two series differed noticeable, the two series agreed well (cc2 = 0.909), with a mean IMET minus SDSU difference and standard deviation of –1.1 ( 3.9%, the IMET reading about 1% lower on average.  Linear regression gave the fit

RHSDSU  = a + b RHIMET,

with a = -5.0 ( 0.3%, b = 1.091 ( 0.011, and a standard deviation of ( 3.9% about the fit (Figure 2).   The IMET relative humidity data were used in subsequent analysis.

Figure 2. Regression of the IMET and SDSU 5-min relative humidity data for

periods when the air temperature was above –4(C.

Barometric Pressure

The IMET and SDSU barometers mounted on the bow mast returned good data for the entire cruise.  After removing some obvious spikes from the SDSU series, the SDSU series were filtered and subsampled to obtain an air pressure series coincident with the 5-min IMET series.  The two series were highly correlated (cc2 = 0.982), with a mean IMET minus SDSU difference and standard deviation of –1.5 ( 0.7 mb, the IMET reading about 1.5 mb lower on average.  Linear regression gave the fit

BPSDSU  = a + b BPIMET,

with a = -3.7 ( 4.7 mb, b = 1.005 ( 0.004, and a standard deviation of ( 0.7 mb about the fit (Figure 3).    The IMET air pressure data were used in subsequent analysis.

Figure 3. Regression of the IMET and SDSU 5-min air pressure data. The two

series track well with several short periods of apparent noise in the

SDSU data.

Incident Short-Wave Radiation

Two IMET units and one ASIMET unit were deployed on Revelle3.  The second IMET unit was placed near the primary IMET unit.  The ASIMET unit was placed on the port side of the O1-deck, about 5 m back from the forward railing, by J. Ware.

IMET unit 1 returned good data for several days before stopping. IMET unit 2 returned only bad (flagged) data.  ASIMET returned good data for entire cruise.

After removing the night-time bias from each record and a few points where the IMET record was constant in time, the IMET and ASIMET were compared using linear regression for the period yd = 16.15 – 19.22 (3.07 days).  The resulting fit was

SWIMET = a + b SWASIMET,

where a = 1.3 ( 0.3, b = 1.035 ( 0.002, with a standard deviation of ~7 W/m2 about the fit. The two series agreed closely, with a cc2 = 0.995.  The IMET read about 3.5% higher than the ASIMET, which is within the manufacturer’s calibration uncertainty. The ASIMET record was used in all analysis.

Incident Long-Wave Radiation

Two IMET units and one ASIMET unit were deployed on Revelle3.  The second IMET unit was placed near the primary IMET unit.  The ASIMET unit was placed on the port side of the O1-deck, about 5 m back from the forward railing, by J. Ware.

All three units returned data for the entire cruise (Figure 4).  The IMET1 and IMET2 records differed by ~50 W/m2, with IMET2 exhibiting much more high frequency variability and an occasional period of constant values. The ASIMET record appeared more similar to IMET2, without as much high frequency variability.  The three records are similar during the center period yd 24–27.5.

To help determine which of these three records is more accurate, the following comparisons were made. First, IMET1 was found to be highly correlated with air temperature, with a cc2 = 0.98 and a standard deviation about the fit of only ~3 W/m2.  Thus, IMET1 was essentially a proxy for air temperature and was not used.  Why all three records tend to track during the period yd 24–27.5 is not clear.

Figure 4. Basic 5-min incident long-wave radiation time series obtained with

the IMET1 (blue), IMET2 (green), and ASIMET (red) units.

The IMET2 and ASIMET records were more similar, with IMET2 having more high frequency variability but slightly weaker lower frequency changes.  Both records exhibit similar timing, with two notable exceptions, the period January 19.95–20.25, when IMET2 appeared stuck, and the period January 27.6–28.3, when the ASIMET record lags IMET2 by about 0.6 days (Figure 5).  After removing these two periods, the remaining IMET2 and ASIMET records were compared using linear analysis, which gave

LWIMET = a + b LWASIMET,

where a = 51.3 ( 1.8 W/m2, b = 0.768 ( 0.012, with a standard deviation of ~11 W/m2 about the fit, and a cc2 = 0.77.  On average, the ASIMET was 8.7 ( 14.0 W/m2 higher than IMET2.  Figure 5 shows that the mean difference between the two records tended to vary on both short and longer time scales on several days, suggesting a sensor or system drift in one or both of the units.  This tendency is shown in the regression plot (Figure 3) by the cluster of values below the regression line.  Both series appear to have a cutoff value near 200 W/m2.

Figure 5. Difference between ASIMET and IMET2 long-wave time series

during Revelle3 cruise.

Figure 6. Regression between the ASIMET and IMET2 long-wave time series

from the Revelle3 cruise.

Pre- and post-cruise comparison tests were made at WHOI between the Revelle3 ASIMET unit (LW205) and the ASIMET unit (LW204) deployed on the winter Khromov cruise.    The post-cruise comparison is split into period A (corresponding to the units as returned from the cruises) and period B (after both units were modified to have two-stage amplification).   During the three comparison periods, the two ASIMET records agreed quite closely with the exception of a mean offset of about 23.7 +/- 2.7 W/m2 prior to the unit modification and 12.6 +/- 2.6 W/m2 after the modification (Table 1).

Table 1. Results of pre- and post-cruise comparison tests for ASIMET LW204

and LW205 units.  The mean and standard deviation of the individual

series and the difference series are given in W/m2.

| Comparison | Period, Sample N | LW204 | LW205 | Difference |
| --- | --- | --- | --- | --- |
| 1999 Pre-Cruise Test | yd = 348.8–355.6 | 246.4 | 270.8 | 24.4 |
|  | N = 9780 | 68.5 | 67.3 | 2.6 |
| 2000 Post-Cruise Test A | yd = 161–192.5 | 318.3 | 341.3 | 23.0 |
|  | N = 45360 | 28.6 | 28.9 | 2.7 |
| 2000 Post-Cruise Test B | yd = 196–318.6 | 356.6 | 369.2 | 12.6 |
|  | N = 176580 | 45.5 | 44.6 | 2.6 |

When the Revelle3 IMET2-ASIMET and the WHOI ASIMET comparisons are considered together, it seems clear that the Revelle ASIMET unit (LW205) was biased high in comparison to both IMET2 and the Khromov ASIMET unit (LW204).  While the calibration and system accuracy of the IMET2 unit is not readily available, the calibration histories of the long-wave sensors used in the two ASIMET units are known to be stable in time, so we chose to use the ASIMET (LW205) record minus a positive bias  as the best measure of downward long-wave radiation during the Revelle3 cruise.  The exact value of  may not be determined from the data available, so we chose = 12 W/m2, one-half the mean bias between LW205 and LW204. (By the same argument, we will add to LW204 to obtain our best estimate of the downward long-wave radiation during the winter Khromov cruise.)  The inherent uncertainties in the WHOI Eppley PIR sensor is roughly ( 1.5-1.8% (or 5 W/m2) and in  is of order ( 10–15 W/m2, giving a rough estimate of the total uncertainty in the downward long-wave radiation of ( 20 W/m2.

Note: An estimate of the net long-wave heat flux into the ocean was made using bulk formula involving measured insolation, RH, and SST.  This approach uses the difference between measured insolation and clear-sky insolation to estimate the cloud cover and cloud correction factor, which is then used to estimate the amount of upward long-wave radiation that is absorbed and re-radiated downward back to the ocean surface.  While the bulk estimate showed occasional similarity with the IMET 2 and ASIMET records, the bulk estimate was not correlated (cc2 < 0.09) with either measured series.   This emphasizes the need for direct measurement of incident long-wave radiation on ships and buoys for air–sea interaction studies.

Sea Surface Temperature

The IMET sea surface temperature sensor, mounted near the bow thruster, returned good data.  At first, the large negative temperature jumps were thought to be in error, possibly due to air and ice being sucked into the intake.  A comparison between the IMET SST data and the temperature measured in the top 4 m with the SeaSoar during the SeaSoar survey work showed excellent agreement.  The 649 comparison values were highly correlated (cc2=0.999), with a mean difference between the IMET and SeaSoar SST of only -0.0001 OC.  The standard deviation of the SST difference   was only +/- 0.0832 OC, due in part because these samples were taken in the surface mixed layer.

Table 2.  Summary of meteorological instrumentation and estimated measurement

uncertainties for Revelle3.

| Variable | Sensor | Uncertainty |  | Notes |
| --- | --- | --- | --- | --- |
| WS | SDSU Anemometer | ___ m/s |  | 1 |
| WD | SDSU Anemometer | ___( |  |  |
| AT | IMET - RM Young | 0.01(C |  | 2 |
| RH | IMET - RM Young | 5% |  |  |
| BP | IMET - RM Young | 2.0 mb |  | 3 |
| SW | ASIMET – Eppley PSP | 3% |  | 4 |
| LW | ASIMET – Eppley PIR | 20 W/m^2 |  | 5 |
| SST | IMET – SeaBird | 0.08(C |  | 6 |

Notes:

1. SDSU anemometers (clive, can you describe…)

2. The IMET AT and RH sensors were mounted at 15 m on the bow mast, and experienced winds above 5 (2.5) m/s 89 (97) % of the time. This coupled with the reduced solar heating due to generally cloudy skies and reduced insolation due to the time of year should minimize additional errors due to solar heating at very low wind speeds.

3. This estimate includes possible dynamical effects due to the airflow over the ship. (Payne, personal communication)

4. Eppley calibration uncertainty.

5. While the ASIMET Eppley PIR sensor is thought to be accurate to roughly 1.5–1.8%, the system uncertainty is larger due in part to the lack of a system comparison with an accepted long-wave system standard (Fairall et al., Payne.. ). This estimate of the uncertainty is based on the inherent sensor uncertainty of 5 W/m2 and one/half the measured bias difference between LW204 and LW205 of 24 W/m2, assuming the true value lies in the middle (see Payne and Anderson, 1999).

6. This uncertainty is based on the standard deviation of the IMET-SeaSoar SST difference series. The mean uncertainty is considerable less, probably of order 0.01 OC. 7. Additional information concerning the ASIMET units is available at http://uop.whoi.edu/uophome

References

Beardsley, R.C., E.P. Dever, S.J. Lentz, and J.P. Dean, 1998. Surface heat flux variability over the northern California shelf. Journal of Geophysical Research, 103(C10), 21,553-21,586.

Fairall, C.W., P.O.G. Persson, E.F. Bradley, R.E. Payne, and S.P. Anderson,

1998. A new look at calibration and use of Eppley precision infrared radiometers. Part I: Theory and Application. Journal of Atmosphere and Oceanic Technology, 15, 1229-1242.

Payne, R.E. and S.P. Anderson, 1999.  A new look at calibration and use of  Eppley precision infrared radiometers. Part II: Calibration and use of the Woods Hole Oceanographic Institution Improved Meteorology Precision Infrared Radiometer.

Journal of Atmosphere and Oceanic Technology, 16, 739-751.

Pawlowicz, R., R. Beardsley, S. Lentz, E. Dever, and A. Anis, 2001. The Air-Sea Toolbox: Boundary-layer parameterization for everyone.  EOS, (in press).